## Slide 1: AN INVESTIGATIONAL STUDY INTO THE DESIGN OF A LOW COST, ADAPTIVE HEARING AID
Presented by:
Kayla-Jade Butkow (714227)
Kelvin da Silva (835842)

## Slide 2: Outline
Objectives and Specifications
System Block Diagram
Simulated vs Hardware Hearing Aid
Hearing Aid Functionality
Cost Breakdown
Results
Future Work and Conclusion

## Slide 3: Objectives and Specifications
To develop a low cost hearing aid
Functionality:
  Amplifying specific frequency bands according to a person’s audiogram
  User tuneable directionality
Done in the form of:
  Software simulation
  Hardware proof of concept

## Slide 4: System Block Diagram

## Slide 5: Simulated vs Hardware Hearing Aid
Property | Simulation | Hardware
Number of microphones | 10 | 4
Device Bandwidth | 0.25–8 kHz | 2.8-3.5 kHz and 5.6-7 kHz
Filter Order | 14 | 2
Filter Bandwidth | 1/3 Octave | 1/3 Octave
Types of filters | Butterworth FIR bandpass | Butterworth bandpass
Number of filters | 16 per microphone | 2 per microphone
Number of steerable angles | 19 (10° increments) | 5 (0°, 60°, 90°, 120°, 180°)
Real time data acquisition | No | Yes

## Slide 6: Compensatory Amplification
Audiogram matching: requires amplification of individual frequency bands
Notes: Independent amplification
  Band-by-band basis
Audiogram matching
  Requires interpolation of audiogram
  Unique for an individual

## Slide 7: Directionality
Amplification in a user specified direction
Delay-and-sum beamforming

## Slide 8: Cost Breakdown
Component | Cost (R)
Arduino Due | 569.00
MAX9814 Microphone Amplifier x4 | 546.00
LM358 x5 | 16.05
LM386 | 15.75
Headphone Jack | 16.31
9V Battery | 70.00
Veroboard | 49.70
Headphones | 69.90
Casing | 59.90
Miscellanous | 50.00
Total | 1462.61

## Slide 9: Testing
Directionality
Compensatory gain

## Slide 10: Results: Simulation
Matched to an audiogram
Average error per frequency band: 1.41%

## Slide 11: Results: Simulation
60°
90°
Most precise steering at 3.15 kHz
Spatial aliasing
  = wavelength of maximum frequency

## Slide 12: Results: Hardware
Input sound frequency: 3.15 kHz
Effects of amplification of one band are seen across the frequency spectrum
  Filter order
Error caused by interaction of stop-bands

## Slide 13: Results: Hardware
90°
60°
Most accurate at 90°
  No time delay
Increased error in other directions
  Integer number of sample shifts
Nulls not distinctive in measurements

## Slide 14: System Error Analysis
Compensatory Amplification | Applied Frequency (kHz) | Frequency Band (kHz) | Error (%)
 | 3.15 | 2.82-3.55 | 0.81
 | 6.30 | 2.82-3.55 | 15.34
 | 3.15 | 5.62-7.08 | 19.56
 | 6.30 | 5.62-7.08 | 3.67
Directionality | Dial Angle (°) | Average Error (%)
 | 0 | 46.6
 | 60 | 30.7
 | 90 | 12.7
 | 120 | 22.7
 | 180 | 51.7
 | Omni-directional | 42.7
Notes: Large omni error due to mics

## Slide 15: Future Work
Higher quality omni-directional microphones
More microphones
  Improve the precision of the beam steering
Integrated circuit chip
  Preprocessing of the audio signals
Emdedding the microphones and circuitry into headphones
  Reduce the size of the device
  Make the device more user friendly

## Slide 16: Conclusion