## Slide 1: Toward the Design of a Low Cost, Adaptive Hearing Aid
Investigators:
Kayla-Jade Butkow and Kelvin da Silva
Presented by:

## Slide 2: Hearing Loss
Affects 5% of the global population
Causes
  Age
  Disease
  Trauma
Decreased quality of life
  Social isolation
  Stigmatization

## Slide 3: The Solution – Hearing Aid
Compensatory amplification
  Correcting an individual’s hearing deficits
  Applied based on an audiogram
Directionality
  Amplifying sounds in particular directions – currently direction of highest SNR
  Attenuating sounds in other directions
Cost of commercial hearing aid ~2000$

## Slide 4: Objectives
Development of a low cost hearing aid
Functionality:
  Amplifying specific frequency bands according to a person’s audiogram
  User tuneable directionality (novelty)
Done in the form of:
  Software simulation
  Hardware proof of concept

## Slide 5: System Block Diagram

## Slide 6: Simulated vs Hardware Hearing Aid
Property | Simulation | Hardware
Number of microphones | 10 | 4
Device Bandwidth | 0.25–8 kHz | 2.8-3.5 kHz and 5.6-7 kHz
Filter Order | 14 | 2
Filter Bandwidth | 1/3 Octave | 1/3 Octave
Types of filters | Butterworth FIR bandpass | Butterworth bandpass
Number of filters | 16 per microphone | 2 per microphone
Number of steerable angles | 19 (10° increments) | 5 (0°, 60°, 90°, 120°, 180°)
Real time data acquisition | No | Yes

## Slide 7: Compensatory Amplification
Audiogram matching: requires amplification of individual frequency bands
Notes: Independent amplification
  Band-by-band basis
Audiogram matching
  Requires interpolation of audiogram
  Unique for an individual

## Slide 8: Directionality
Amplification in a user specified direction
Delay-and-sum beamforming
Source

## Slide 9: Testing
Directionality
Compensatory gain

## Slide 10: Simulation Results: Amplification
Matched to an audiogram
Average error per frequency band: 1.41%

## Slide 11: Simulation Results: Directionality
60°
90°
Most precise steering at 3.15 kHz
Spatial aliasing
  = 5cm
  = wavelength of maximum frequency (3.4 kHz)

## Slide 12: Hardware Results: Amplification
Total cost: $100
Input sound frequency: 3.15 kHz
Error caused by interaction of stop-bands
Notes: Effect of amplifying one band is seen across the frequency spectrum- due to the order of the filter

## Slide 13: Hardware Results: Directionality
90°
60°
Most accurate at 90°
  No time delay
Increased error in other directions
  Integer number of sample shifts

## Slide 14: System Error Analysis
Compensatory Amplification | Applied Frequency (kHz) | Frequency Band (kHz) | Error (%)
 | 3.15 | 2.82 - 3.55 | 0.81
 | 6.30 | 2.82 - 3.55 | 15.34
 | 3.15 | 5.62 - 7.08 | 19.56
 | 6.30 | 5.62 - 7.08 | 3.67
Directionality | Dial Angle (°) | Average Error (%)
 | 0 | 46.6
 | 60 | 30.7
 | 90 | 12.7
 | 120 | 22.7
 | 180 | 51.7
 | Omni-directional | 42.7
Notes: Large omni error due to mics

## Slide 15: Future Work
Higher quality omni-directional microphones
Integrated circuit chip
  Pre-processing of the audio signals
Embedding circuitry into headphones
  Reduce the size of the device
  Make the device more user friendly

## Slide 16: Conclusion
Objectives and specifications have been met
Low cost – $100
Full hearing aid simulation
  Compensatory amplification
  Steerable directionality
Concepts proven in hardware

## Slide 17: References
[1] L. Tiete et al. “Detecting Laterality and Nasality in Speech with the Use of a Multi-Channel recorder.” Sensors (Basel, Switzerland), vol. 14, pp. 1918-1949, 02 2014.

## Slide 18: Questions?